## Sheet1
Portable Ventilator MR Compatible - Quantity 1
Category | Specification
Adult and Pediatric patients . Specify
Magnetic Field Compatibility
Certified for 1.5T or 3T MR environments; safe operation with no gauss line restriction.
Portability
Compact and lightweight design for easy transport.
System type
Specify if system includes batteries
Ventilation Modes
Volume-Controlled Ventilation (VCV), Pressure-Controlled Ventilation (PCV).
Pressure Support Ventilation (PSV), SIMV, and CPAP.
Optional: Adaptive or Non-Invasive Ventilation (NIV).
Performance Parameters
Tidal Volume Range: Adjustable to 1500 mL.
Respiratory Rate: Adjustable from 5 to 60 breaths per minute.
PEEP: Adjustable from 0 to 20 cmH₂O.
Inspiratory Flow Rate: Up to 180 L/min.
Monitoring and Alarms
Displays: Airway pressure, tidal volume, respiratory rate, FiO₂ levels.
Alarms: High/low pressure, disconnection, power failure.
Safety Features
Emergency shut-off button.
Accessories
MRI-compatible breathing circuits. Single use.
Disposable filters.
Mounting options for MRI suites.
Regulatory and Certification
CE or FDA certified for MR compatibility.
Compliance with IEC 60601-2-33 or equivalent MR safety standards.
Warranty
Delivery